[image: image1.jpg]


[image: image2.jpg]


[image: image3.png]Ordering Information

Desciption Waight

bsisgin
o Tabletop

‘Vbraamped seelwhit plastclaminate 30

VibraDamped see, .

mognetc anlessstel loinate
Vrabamped sed backpasticlaminate
142 1"VibraDamped seel, »
it st e
T Gray Granite -

s





[image: image4.jpg]TRANSHSSBLTY

HISOTON



[image: image5.jpg]THBLETOP, -Gyl

- loPTionAL)

LevEnG
HES




          

[image: image6.jpg]PaDDED.

o
o

TN

7

o=

o]

LonG
ity
oFtonay

ns







BRIDGE


Technology


Semiconductor Equipment Sales











Bus: (480) 219-9007    Fax: (480) 452-0172


Email:  sales@bridgetec.com


Web:  http://www.bridgetec.com/








1200 Series�LabMate�VIBRAPLANE Workstations





Features / Benefits� �•  Suitable for moving loads�   and higher centers of gravity�•  Ultra low vertical�   natural frequencies�•  Isolation efficiencies�   approaching 99%�•  High performance�   Active-Air Suspension�•  Automatic leveling and�   height control�•  VibraDamped for�   improved performance�•  1" Thick VibraDamped Steel�   or 4" granite tabletops�•  Choose from 16 sizes�•  Class 100 Cleanroom Compatible�   - Class 10 available�  �The 1200 Series LabMate Workstation is a proven performer providing a reliable vibration-free work environment..�  �Specifications 


�


Min. Load @ 20 psi� 


Vertical Natural Frequency�
2.1Hz�
�
   Isolation Efficiency @  5Hz�
43%�
�
   Isolation Efficiency @ 10Hz�
89%�
�
�


Horizontal Natural Frequency�
3.5Hz�
�
   Isolation Efficiency @  5Hz�
40%�
�
   Isolation Efficiency @ 10Hz�
80%�
�
�


Max. Load @ 80 psi� 


Vertical Natural Frequency�
1.6Hz�
�
   Isolation Efficiency @  5Hz�
76%�
�
   Isolation Efficiency @ 10Hz�
93%�
�
�


Horizontal Natural Frequency�
2.5Hz�
�
   Isolation Efficiency @  5Hz�
75%�
�
   Isolation Efficiency @ 10Hz�
88%�
�
�


Gross Load Capacity @ 80 psi�   1201   1300 lbs.(590kg)�   1202   2800 lbs.(1270kg) 


�


Finish:�standard   Charcoal Grey Medium�standard   Textured Polyurethane�Class 10   White Epoxy Powder Coat 


�


Note: Typical performance efficiencies are for microdisturbances � � �Applications�(Partial List)� �Surface Plates, Analytical Balances, Cell Injection, Confocal Microscopes, Electro Physiology, Optical Microscopes, Tunneling Microscopes, Wafer Probing, Roundness Checkers, Mask Aligners, and CMMs. 








